## Slide 1: Ion temperature imaging in the divertor of the W7-X stellarator using coherence imaging spectroscopy
Matt Kriete1
V. Perseo2, D. Gradic2, D.A. Ennis1, R. König2, D.A. Maurer1, and the W7-X Team2
1Auburn University
2Max Planck Institute for Plasma Physics
High Temperature Plasma Diagnostics Conference
April 22, 2024
This work has been funded by US DoE grant DE-SC0014529. This work has been carried out within the framework of the EUROfusion Consortium, funded by the European Union via the Euratom Research and Training Programme (Grant Agreement No 101052200 -- EUROfusion). Views and opinions expressed are however those of the author(s) only and do not necessarily reflect those of the European Union or the European Commission. Neither the European Union nor the European Commission can be held responsible for them.

## Slide 2: The coherence imaging spectroscopy technique has been advanced to image impurity ion temperatures
Coherence imaging spectroscopy (CIS) provides 2D images of plasma parameters in the core and edge of magnetically confined plasmas
Flow velocity measurements now routinely available with CIS
New: we have advanced the CIS technique to reliably measure ion temperatures
  This talk: impurity in edge plasma
  Wednesday talk by R. Lopez Cansino [3.1.2]: main ion in core plasma
C2+ temperature distribution near lower divertor in W7-X

## Slide 3: Outline
Motivation and overview of coherence imaging spectroscopy
CIS design principles for edge measurements
Inference of from CIS data
Validation of CIS measurements on W7-X

## Slide 4: Outline
Motivation and overview of coherence imaging spectroscopy
CIS design principles for edge measurements
Inference of from CIS data
Validation of CIS measurements on W7-X

## Slide 5: Images of edge ion temperature provide crucial information for understanding stellarator divertor physics
Many key scrape-off layer (SOL) plasma parameters depend on
  information is sparser than other parameters (, , )
Stellarators SOLs are fully 3D → diagnostics with high spatial information are needed to fully characterize edge plasma
Physical sputtering yield strongly depends on divertor ion temperature
P.C. Stangeby, PPFC 60 044022 (2018)
0°
72°
W7-X SOL consists of magnetic islands intersected by divertor targets
Notes:   Can mention a few edge plasma parameters that depend on Ti: divertor erosion rates, ion sound speed, sheath heat transmission factor

## Slide 6: Coherence imaging spectroscopy is an established diagnostic for impurity emission and velocity measurements
Coherence imaging spectrometer (CIS) compared to dispersive spectrometer:
  CIS obtains substantially more spatial information
  CIS has ~100x higher optical throughput (no slit required)
  CIS collects substantially less spectral information → requires isolated spectral line
V. Perseo et al, NF 59 124003 (2019)

## Slide 7: CIS encodes Doppler information from spectral emission lines into a spatial interference pattern
Doppler shift
Ion velocity shifts wavelength of the emission line, resulting in a shift of the interference pattern
demodulation + calibration
Ion velocity (processed data)
Interferogram (raw data)
Doppler broadening
Ion temperature broadens the emission spectrum, resulting in reduced contrast
[Figure courtesy N. Allen]

## Slide 8: Interference pattern formed by imaging plasma light through a polarization interferometer
imaging fiber bundle
V. Perseo et al, RSI 91 013501 (2020)
: intensity
: phase
: contrast
Polarizers + birefringent crystal split light into two orthogonally polarized beams and introduce phase delay between them
Instrument sensitivity given by group delay:
Notes: Rework this to use animations to break this up (and include pictures of actual hardware is shown in addition to schematics + maybe polarization state of light at each stage?)

## Slide 9: Outline
Motivation and overview of coherence imaging spectroscopy
CIS design principles for edge measurements
Inference of from CIS data
Validation of CIS measurements on W7-X

## Slide 10: Main challenge measuring edge with CIS is distinguishing between Doppler broadening and Zeeman splitting
Several mechanisms can contribute to the CIS contrast
  Doppler broadening
  Zeeman splitting : depends on magnetic field and observation angle
  Stark broadening
  Background emission (e.g. bremsstrahlung)
  Instrument contrast : calibration factor
Choice of C III 465 nm spectral line greatly simplifies analysis
  C III intensity >> background
  Well-separated from other strong impurity lines
  Inferred is decent proxy for main ion up to ~30 eV
and are comparable in the SOL plasma of magnetic confinement fusion experiments
B = 0 T
Ti = 10 eV
B = 2.2 T
Ti = 10 eV
C III 465 nm line shape
D. Gradic et al, NF 61 106041 (2021)

## Slide 11: CIS birefringent crystals are optimized to minimize sensitivity to Zeeman splitting of C III 465 nm line
Sensitivity to multiplet + Zeeman splitting has non-monotonic group delay dependence → certain group delays give low Zeeman splitting sensitivity
Response to Doppler broadening ()
Response to Zeeman splitting ( & field-sightline angle)
Relative sensitivity to Doppler vs Zeeman
Sensitivity to Doppler broadening increases monotonically with increasing group delay
Specific group delays give higher sensitivity to than in the relevant 2–3 T range
2620 waves chosen

## Slide 12: Multi-delay CIS approach increases the available spectral information, improving measurement accuracy
Multi-delay CIS approach1 obtains spectral information at 4x as many delays as conventional CIS
1J.S. Allcock et al., RSI 92 073506 (2021)
Micropolarizer grid
Polarizer #1
Polarizer #2
Crystal #1 (displacer)
Crystal #2 (delay)
Filter wheel
Collimating lens
Camera lens
Camera with micropolarizer sensor
Quarter-wave plate
Multi-delay fringe pattern
Possible solutions for and (1% uncertainty in contrast measurement)
solution space using information from all four group delays
Combination of linear carrier (crystal #1) and pixelated carrier (crystal #2) interferometry
Polarizers plus crystal #1 encode spectra into linear fringes at
Crystal #2, QWP, and polarization camera encode spectra into pixelated fringes at
Fringes multiplied together → information also encoded at and

## Slide 13: Multi-delay CIS instrument characterized and calibrated using precision tunable laser
Group delay measured using fine wavelength scans of precision tunable C-WAVE laser1,2
1D. Gradic et al., FED 146 995 (2019)
2V. Perseo et al., RSI 91 013501 (2020)
All group delays within 5% of design value
Variation of group delay across image measured to eliminate errors caused by common assumption of constant
For more details on laser scans, see posters by V. Perseo [1.2.30] and S. Ahkundzada [1.4.26]
Notes: Sapida Ahkundzada

## Slide 14: Outline
Motivation and overview of coherence imaging spectroscopy
CIS design principles for edge measurements
Inference of from CIS data
Validation of CIS measurements on W7-X

## Slide 15: Magnetic field geometry is precisely known along CIS sightlines
toroidal direction
W7-X stellarator optimized for low plasma current → magnetic field generated largely be external coils and easy to calculate along CIS sightlines

## Slide 16: Temperature is inferred by fitting a model of the C III line shape to data, using variation of magnetic field along sightlines
Data: phase and contrast images for each of the four group delays
Fitting model: C III 465 nm triplet with Doppler shift/broadening and Zeeman splitting
C III emission localized to a single 1-5 cm long region in SOL near divertor1,2 → and are not independent fit parameters
Fit parameters:
C2+ velocity
C2+ temperature
Distance of emission location along sightline
Magnetic field known along sightlines → ,
1D. Gradic et al., PPCF 64 075010 (2022)
2M. Krychowiak et al., EPS 2021 P1.1026
Contrast
Phase
CIS viewing direction
Tomographic reconstructions of C III radiation distribution2
and for example sightlines

## Slide 17: Outline
Motivation and overview of coherence imaging spectroscopy
CIS design principles for edge measurements
Inference of from CIS data
Validation of CIS measurements on W7-X

## Slide 18: -optimized CIS diagnostic is validated by cross-comparison with a high-resolution dispersive spectrometer
CIS & spectrometer data averaged over a 60 s stationary time window → high signal-to-noise ratio
Direct comparison of CIS and spectrometer data requires Fourier transformation of spectrometer wavelength spectra into coherence spectra:
spectrometer fiber locations
EJM007+2520
CIS C III intensity image in long-pulse W7-X plasma
spectrometer fiber locations
Fourier transform
Contrast =
: coherence
: interferometer group delay
: center wavelength of line
: intensity per unit wavelength
: integrated line intensity
Spectrometer sightline 38
Spectrometer sightline 38

## Slide 19: CIS and spectrometer measure same spectral information and fits are in good agreement for most sightlines
Highest signal
Most spectral contamination
Lowest spectrometer signal
triplet components weighted 15% higher than statistical values due to strong electron impact excitation [J.D. Hey et al., J. Phys. B 35 1525 (2002)]
spectrometer fiber locations

## Slide 20: Comparison between CIS and Sopra shows excellent agreement of C2+ temperatures
spectrometer fiber locations
Similar C III intensity profiles (difference in vignetting by plasma facing lens)
CIS & Sopra C2+ show excellent agreement:
  Well-matched profile shape
  ~10% difference on average
  25% difference in worst case
Do CIS and the spectrometer infer the same plasma parameters?
horizontal divertor target
pumping gap
vertical divertor target
C III intensity comparison between CIS & spectrometer
C2+ temperature comparison between CIS & spectrometer

## Slide 21: Divertor ion temperature image shows toroidally elongated structures and inverse correlation between and C III radiation
C2+ ion temperature image
over divertor
Regions of high C III emission have lower and vice-versa
Band with particular high (up to 40 eV)
C III intensity

## Slide 22: Summary and outlook
CIS diagnostic technique has been advanced to image C2+ ion temperatures without requiring cross-calibration by a separate dispersive spectrometer
Temperature measurements enabled by multi-delay configuration and novel approach to designing CIS birefringent crystals that minimizes instrument sensitivity to Zeeman splitting
CIS C2+ measurements validated by comparison against high-resolution spectroscopy in W7-X
New measurement capability is being used to characterize SOL distribution across W7-X parameter space (example: density ramp to detachment)
20230214.042
C III intensity
C2+ temperature
Density ramp to detachment